CLRC Instrumentation Department

Problem Report


Problem Report 



 Report Number :- 06

Project Name:-CMS-FED Back end

Item:-




PC 3025M Mk1

Problem:

Enable pin (1) on U135 TPS75725 is connected to 3.3V. 

This caused the linear regulator to be disabled. 



Originator (sign, date)

Project Manager(sign, date)

Remedial Action:

Was rectified by connecting the pin to ground.

Modify circuit schematics for CMS FED V2. The symbol should be updated and the circle replaced with a negation bar



Project Manager (sign, date)



Last modified 24/06/03 10:13

page 1 of 1

