CLRC Instrumentation Department

Problem Report


Problem Report 



 Report Number :- 20

Project Name:-CMS-FED Backend

Item:-




PC 3025M Mk1

Problem:

It is possible to short the board supply via the JTAG connector. 

It is possible to accidentally short the board by miss connecting the JTAG connector.



Originator (sign, date)

Project Manager(sign, date)

Remedial Action:

It is suggested that a proper 2mm or 0.1in pitch connector is fitted, and that the Datamate series such as used on the signal injector card used

Project Manager (sign, date)



Last modified 24/06/03 10:39

page 1 of 1

