CLRC Instrumentation Department

Problem Report


Problem Report 



 Report Number :- 26

Project Name:-CMS-FED Back end

Item:-




PC 3025M Mk1

Problem:

Symbol for crystal is incorrect 

The symbol for the crystal U265 is incorrect. It does not show the shutdown pin



Originator (sign, date)

Project Manager(sign, date)

Remedial Action:

Discuss with DO before next production run.

Has been revised



Project Manager (sign, date)



Last modified 24/06/03 14:51

page 1 of 1

