CLRC Instrumentation Department

Problem Report


Problem Report 



 Report Number :- 29

Project Name:-CMS-FED Back end

Item:-




PC 3025M Mk1

Problem:

12V switch can't run at full load 

The ripple caused by the action of the 12V to -5V switcher is causing the MAX5906 part to trip at start-up. 



Originator (sign, date)

Project Manager(sign, date)

Remedial Action:

Suggest fitting a filter such as a 15uH choke to reduce the supply current ripple. This requires further investigation with a 50MHz bandwidth probe.

The switcher has been changed and the new part has independent ground planes and substantial filtering.

Project Manager (sign, date)



Last modified 24/06/03 14:54

page 1 of 1

