CLRC Instrumentation Department

Problem Report


Problem Report 



 Report Number :- 37

Project Name:-CMS-FED back end

Item:-




PC 3025M Mk1

Problem: TTCRX does not properly reset on short power cycles.

The TTCRX does not properly reset if power is briefly interrupted. It is supected that the voltage on the power planes does not drop low enough to initate a proper reset. The reset to the TTCRXBGA is to be checked.

Originator (sign, date)

Project Manager(sign, date)

Remedial Action:



Project Manager (sign, date)



Last modified 10/09/03 11:24

page 1 of 1

