CLRC Instrumentation Department

Problem Report


Problem Report 



 Report Number :- 42

Project Name:-CMS-FED back end

Item:-




PC 3025M Mk1

Problem: Pre-pulse undershoot of AD9218 in 2Vpp mode.

There is a pre-pulse undershoot (5% of pulse height) observed on AD9218 is operated in 2V pp mode.

Problem has also been observed on AD9218 evaluation board. However it is not seen in 1 V pp mode.

Analog Devices technical experts are investigating.

Originator (sign, date)     John Coughlan     24/10/03

Project Manager(sign, date)    John Coughlan     24/10/03

Remedial Action:

Operate AD9218 in 1 V pp mode and reduce gain in analogue chain by 2 by either:

1. Change load Resistor on OptoRx output from 100 Ohm to 50 Ohm.

2. Add 100 Ohm in parallel across output of Elantec Opamp.



Project Manager (sign, date)



Last modified 24/10/03 12:43

page 1 of 1

