CLRC Technology Department

Procedure 9010 Form 1.0

Route Card
Project

CMS FED
Product

Front End Driver
Identification Ser. No. PC3205M/1      001



manufacturing process / supplier
Batch ID / date

PCB manufactured by Express Circuits Ltd
6 off December 2002

Assembled by SAETECH
2 off January 2003




operation
Comments
initial
date

Problem Notes From SAETECH company
Problem report from the company outlining component assembly issues.
SAETECH
22/01/03

Visual Inspection 
Noted problems with footprints for:

Diodes HBAT540 and 47CTQ020S.
JS, ITC
23/01/03

Pre power-up measurements
See document

Power Checkout Document
JS, ITC
23/01/03

Board Power up failed
Removed C3402.
JS,

ITC
24/01/03

Board Power up failed
Changed R1240 & R394 to 470K
JS, ST
27/01/03

Cut LK59 and LK60 and swap the signals
See PR003
ITC
23/01/03

Remove U137, R373, R372
See PR003
JS, ST
27/01/03

Lift pin 1 of U135 and connect to pin 3 of U135
See PR006
JS, ST
27/01/03

Add 100nf cap in parallel with R388
See PR010
JS, ST
27/01/03

Lift pin 8 of U128

JS, ST
27/01/03

Attach serial number
Yes
ST
27/01/03

Record input supply voltages
See document

Supplementary Document
JS
28/01/03

Detect JTAG chains
JTAG1 [ Pass]

JTAG2 [ Pass]

JTAG3 [ Pass]
RM, ST
29/01/03

Test Chains with no loop back for external signals
JTAG1 [ Pass]

JTAG2 [ Pass]

JTAG3 [ Pass]
RM, ST
30/01/03

Power recycling due to under voltage trip on 3.3V rail.
Remove R400 & R401. Connect pin 9 to pin 14 on U144.
ITC
06/02/03

Test Chains with loop back for external signals
JTAG1 [ Pass]

JTAG2 [ Pass]

JTAG3 [ Pass]
RM, EJF,

ITC
06/02/03

Fitted two opto devices

ITC















































PAGE  
2

