CLRC Technology Department

Procedure 9010 Form 1.0

Route Card
Project

CMS FED
Product

Front End Driver
Identification Ser. No. PC3205M/2      018



manufacturing process / supplier
Batch ID / date

PCB manufactured by DDi
2 off July 2004

Assembled by DDi
2 off July 2004




operation
Comments
initial
date

Visual Inspection
Failed,
Board made to quotation version not final version.

R1174 & R1178 missing.

C27 & C31 not fitted (incorrect ones sent), now fitted.

D1_1 etc (all HBAT-540Cs) fitted with wrong part.
ITC

ITC
03/08/04

04/08/04

JTAG tests
Failed, 

TTC chip returns incorrect ident and fails all tests.

QDR memories fail all tests. Fixed by connecting TRST line high instead of low, can be fixed by setting the TRST line in the JTAG software.
VME device fails all tests,  fixed, four dry joints found on PROM.
Eight solder bridges on series ADC resistors. Fixed
ITC

ITC
03/08/04

04/08/04


Wired-up front panel switch
ITC
14/09/04

FED Acceptance test
Passed
ITC
23/02/05

Opto serial numbers
U7_1    380264

U7_2    380262

U7_3    350446

U7_4    380266

U7_5    350348

U7_6    380268

U7_7    380267

U7_8    350447
ITC
23/02/05
















PAGE  
1

