CLRC Technology Department

Procedure 9010 Form 1.0

Route Card
	Project

CMS FED
	Product

Front End Driver
	Identification Ser. No. PC3205M/2      020


	manufacturing process / supplier
	Batch ID / date

	PCB manufactured by DDi
	2 off July 2004

	Assembled by DDi
	2 off July 2004

	
	

	operation
	Comments
	initial
	date

	Visual Inspection
	Failed,

Board made to quotation version not final version.

C27 & C31 not fitted (incorrect ones sent), now fitted.

D1_1 etc. (all HBAT-540Cs) fitted with wrong part.
	ITC

ITC
	03/08/04

04/08/04

	JTAG tests
	Failed, 

TTC chip returns incorrect ident and fails all tests.

QDR memories fail all tests. Fixed by connecting TRST line high instead of low, can be fixed by setting the TRST line in the JTAG software.
Four solder bridges on series ADC resistors. Fixed
	ITC
	03/08/04

	
	DS2_3 faulty, Replaced.
	ITC
	14/09/04

	
	Wired-up front panel switch
	ITC
	14/09/04

	FED acceptance test
	Passed
	ITC
	03/03/05

	Opto serial numbers
	U7_1    2X0135

U7_2    2X0134

U7_3    2X0133

U7_4    2X0139

U7_5    2X0136

U7_6    2X0137

U7_7    2X0138

U7_8    2X0140
	ITC
	03/03/05

	Returned Faulty
	TTC problem (fibre?)
	ITC
	30/03/07

	
	
	
	


PAGE  
1

